## Direct POCT
WHO TPP: Point of care test (POCT) for suspected COVID-19 cases and their close contacts to diagnose acute SARS-CoV-2 infection in areas where
reference assay testing is unavailable, or turnaround times obviate utility.
Published: 28 Sept 2020
| Wolrd Health Organization v.1.0 (20 Sept 2020) |  | FDA EUA Selected Direct (antigen/RNA) POCT (performance based on manufacturer's claims) | [6 empty] | FIND Selected Direct (antigen/RNA) POCT (performance based on independent evaluation)
| Acceptable | Desirable | BinaxNOW COVID-19 Ag Card | Sofia 2 Flu+SARS Antigen FIA | LumiraDx SARS-CoV2 Ag Test | COVID-19 Home Test | MobileDetect Bio BCC19 Test (RT-LAMP) | AQ-TOP Plus COVID-19 Rapid Test (RT-LAMP) | SARS-CoV-2 Detectr Test (RT-LAMP) | Bioeasy Diagnostic kit - COVID-19 Ag test | Sofia 2 SARS Antigen FIA | Standard Q Covid-19 Ag FIA
Manufacturer |  |  | Abbott | Quidel Corporation | LumiraDx | Ellume | DetectaChem Inc. | SEASUN BIOMATERIALS, Inc. | Mammoth Biosciences, Inc. | Shenzhen Bioeasy Biotechnology Co.,Ltd. | Quidel Corporation | SD Biosensor, INC.
Country |  |  | USA | USA | UK / USA | Australia | USA | Korea | USA | China | USA | Korea
URL |  |  | https://dam.abbott.com/en-us/homepage/coronavirus/BinaxNOW-Fact-Sheet-Final.pdf | https://www.quidel.com/immunoassays/sofia-2-flu-sars-antigen-fia | https://www.lumiradx.com/us-en/what-we-do/diagnostics/test-technology/antigen-test | https://www.ellumehealth.com | https://mobiledetectbio.com/ | http://www.seasunbio.com/english/products4.html | https://mammoth.bio/covid/ | http://en.bioeasy.com/?p=1038 | https://www.quidel.com/immunoassays/sofia-2-flu-sars-antigen-fia | http://sdbiosensor.com/xe/product/7672
Intended Use | 1) In areas with confirmed SARS-CoV-2 community wide transmission or confirmed outbreaks in closed or semi-closed communities and in high risk groups; 2) In suspected SARS-CoV-2 outbreak situations; 3) Monitor trends in disease incidence |  | qualitative detection of SARS-CoV-2 antigen in patient care setting; POCT | qualitative detection of SARS-CoV-2 antigen in patient care setting; POCT | qualitative detection of SARS-CoV-2 antigen in patient care setting; POCT | qualitative detection of SARS-CoV-2 antigen; non-prescription home use | qualitative detection of SARS-CoV-2 RNA | qualitative detection of SARS-CoV-2 RNA | qualitative detection of SARS-CoV-2 RNA | qualitative detection of SARS-CoV-2 in patient care setting; POCT | qualitative detection of SARS-CoV-2 in patient care setting; POCT | qualitative detection of SARS-CoV-2 in patient care setting; POCT
Target Population / patient | patients with respiratory symptoms or fever or other symptoms (diarrhoea, anosmia) and either a known contact with a confirmed or probable COVID-19 patient OR living in an area of cluster OR community transmission, and close contact of index patients (confirmed COVID-19 patients) |  | individuals suspected of COVID-19 within 7 days of sympton onset | individuals suspected of COVID-19 within 5 days of sympton onset | individuals suspected of COVID-19 within 12 days of sympton onset | individuals suspected of COVID-19 | individual suspected of COVID-19 | individual suspected of COVID-19 | individual suspected of COVID-19 | individual suspected of COVID-19 | individuals suspected of COVID-19 within 5 days of sympton onset | post
Target use setting | outside laboratories including routien screening points of health care facilities by health care workers or lab techs with appropirate training to use the test. | same as acceptable but can be self-administered and/or performed by trained lay workers (volunteer, community health workers) | patient care setting; CLIA certified labs | patient care setting; CLIA certified labs | patient care setting; CLIA certified labs | non-prescription home use | certified laboratories | certified laboratories | certified laboratories | not specified (patient care setting / labs) | patient care setting; CLIA certified labs | not specified (patient care setting / labs)
Target molecule | SARS-CoV biomarker (RNA, antigen) specific for acute (first week after onset of symptoms/current infection) | SARS-CoV-2 only biomarker specific for acute and subacute (first 2 weeks after onset of symptoms/current infection) | SARS-CoV-2 nucleocapsid protein (and SARS-CoV) | SARS-CoV-2 nucleocapsid protein | SARS-CoV-2 nucleocapsid protein | SARS-CoV-2 & SARS-CoV nucleocapsid protein | SARS-CoV-2 N and E genes | SARS-CoV-2 ORF1ab and N genes | SARS-CoV-2 N genes | n/a | SARS-CoV-2 nucleocapsid protein | SARS-CoV-2 nucleocapsid protein
Analytical sensitivity / LOD | ~ 106 genomic copies/ml | ~ 104 genomic copies/ml | 22.5 TCID50/swab (1125 TCID50/ml) | 113 TCID50/ml | 32 TCID50/ml | 6310 TCID50/ml | 75 copies/µl | 1 copies/µl | 20 copies/µl | n/a | 113 TCID50/ml | 494.5 TCID50/ml
Analytical specificity | detects all SARS-CoV-2 viral strains and dose not cross react with other human coronaviruses (except SARS-CoV-1) or any other common diseases with similar symptoms (influenza A, B, RSV, malaria, dengue) | detects all SARS-CoV-2 virla strains and dose not cross react with other human CoV or common diseases with similar symptoms | no cross-reactivity in a panel of high prevalence respiratory pathogens | no cross-reactivity in a panel of high prevalence respiratory pathogens | no cross-reactivity in a panel of high prevalence respiratory pathogens | no cross-reactivity in a panel of high prevalence respiratory pathogens | no cross-reactivity in a panel of high prevalence respiratory pathogens | no cross-reactivity in a panel of high prevalence respiratory pathogens | no cross-reactivity in a panel of high prevalence respiratory pathogens | n/a | no cross-reactivity in a panel of high prevalence respiratory pathogens | n/a
Sensitivity (or PPA) | ≥ 80% | ≥ 90% | 97.1% (85.1 - 99.9) | 96.7( (83.3-99.4) | 97.6% (91.6-99.3) | 95% (82-99) | 97.7% (88-100) | 100% (91-100) | 95% (83.5.98.6) | 93.9% (86.5-97.4) | 93.8% (79.2-99.2) | 90.6% (75-98)
Specificity (or NPA) | ≥ 97% | ≥ 99% | 98.5% (92.0-100%) | 100% (97.9-100) | 96.6% (92.7-98.4) | 97% (93-99) | 100% (94.3-100) | 100% (94-100) | 100% (94.2-100) | 100% (92.1-100) | 96.9% (83.8-99.9) | 96.9% (83.8-99.9)
Type of analysis | qualitative (yes/no), semi-quantitative, or quantitative | N/A | lateral flow immunoassay | lateral flow immunofluorescent | microfluidic immunofluorescence | lateral flow immunofluorescent | RT-LAMP | RT-LAMP | CRISPR-based RT-LAMP | fluorescent immunochromatographic, double antibody sandwich rxn | lateral flow immunofluorescent | fluorescent immunoassay
Interpretation | visual manual and/or hardware reader (proprietary or smart phone) | visual manual read or digital readout via smartphone app reader with connectivity | visual manual read | digital readout | digital readout | digital readout (smartphone) | visual color recognition | digital readout | digital readout | digital reader, fluorescence analyzer | digital readout | digital readout
Sample type | nasopharyngeal, oropharyngeal, nasal swab, nasal wash, sputum | anterior nares (swab), saliva/oral fluid, sputum | nasal swab | nasal swab | nasal swab | nasal swab | nasal swab | nasal swab; BAL | nasal swab / wash | nasal swab | nasal swab | nasal swab
Sample collection device | compatible with an existing swab material | compatible with multiple swab materials (self-collection or no swab e.g. saliva) | swabs in the kit ONLY | swabs in the kit ONLY | recommended swab (not included) | swabs in the kit ONLY | recommended swab (not included) | recommended swab (not included) | recommended swab (not included) | not provided | swabs in the kit ONLY | swab provided in kit
End use profile | trained staff in health care facilities | trained staff or community level (lay person) or self-administered | medical professional or trained staff in certified patient care setting | trained clinical laboratory personnel and individuals trained in point of care settings | trained clinical laboratory personnel and individuals trained in point of care settings | home use | trained lab personnel | trained lab personnel | trained lab personnel | trained medical / lab operator | trained clinical laboratory personnel and individuals trained in point of care settings | professional use
Training needs | 0.5 days with instructions for use and quick reference guide(s) | 2hr with instructions, quick reference guide including smart phone app(s) | N/A (procedure card included) | N/A | N/A | no | Yes | Yes | Yes | n/a | N/A | n/a
Sample preparation steps | 1 | 0 | 1 | bar: N/A to n/a
Reagents reconsitituion | very simple reconstituion | ready to use | ready to use | ready to use | ready to use | ready to use | ready to use | ready to use | ready to use | ready to use | ready to use | ready to use
Sample min volume | single swab and minimal extraction buffer/diluent |  | single swab + extraction buffer | single swab + extraction buffer | single swab + extraction buffer | single swab + extraction buffer | single swab + extraction buffer | single swab + extraction buffer | single swab + extraction buffer | single swab + extraction buffer | single swab + extraction buffer | single swab + extraction buffer
Need to transfer a precise volume of sample | autofill or graduated volume markings on sample transfer device | no, or limited to a number of drops | no, few drops | no, few drops | no, few drops | no, few drops | pipette fixed volumes | pipette fixed volumes | pipette fixed volumes | no, few drops | no, few drops | no, few drops
Number of timed steps | ≤ 3 | 1, with the potential for digitally guided workflows and built-in timers to reduce user errors on timed steps | 1 | 1 | 1 | 1 | 2 | 2 | 2 | 1 (automatic by analyzer) | 1 | 1
Time to result | ≤ 40 min | ≤ 20 min | 15min | 15min | 12min | 15min | 30min | 30min | 45min | 10min | 15min | 30min
Sample stability pre-testing | 30min (10-35oC); 2-4hr (4-8oC); several days frozen (min -20oC) | test compatible with both dry and preserved samples; 3hr (10-40oC); 24hr (4-8oC); months frozen | <1hr room temp | 48hr room temperature or 2-8oC | 5hr room temp; 5 days frozen -80oC | n/a | n/a | n/a | n/a | <24hr at 2-8OC; or -70oC long-term | 48hr room temperature or 2-8oC | 24h room temperature; 2-8oC for 48hr
Result validity stability | fixed reading time | stored image or 6 weeks | <30min | <30min | n/a | n/a | n/a | n/a | n/a | n/a | <30min | n/a
Invalid rate | <2% invalid results with correct use by operator | <0.5% invalid results with correct use by operator | n/a | n/a | n/a | 8% | n/a | n/a | n/a | n/a | n/a | n/a
Operating conditions | 15-35oC, 25-80% RH up to 1500m | 10-40oC; 25-90% RH up to 3000m | Room temperature | room temperature | 10-30 oC; RH 10-90% | room temperature | room temperature | room temperature | room temperature | 2-35oC; RH 40-90% | room temperature | room temperature
Test kit stability and storage | 12mo 4-30oC; tolerates brief period >40oC; RH 75% | 18-24mo at 4-40oC; tolerates freezing and brief period >40oC; RH 75% | 2-30oC; expiration n/a | room temperature, 15-30oC | 2-8oC | 2-30oC | freeze at -20oC | freeze at -20oC | freeze at -20oC | 2-30 oC | room temperature, 15-30oC | room temperture (2-30oC)
Stability of kit once opened | 30min | 1hr | n/a (use immediately) | 12hr | 30 day | 60min | n/a | n/a | n/a | 60min | 12hr | n/a
Specimen capacity & throughput | >5/hr per operator | >10/hr per operator | n/a | n/a | n/a | n/a | n/a | n/a | 1500 sample in 8hr | n/a | n/a | n/a
Safety precautions | standard respiratory sample collection safety precautions recommended; | tests that minimize the need for biosafety requirement are strongly preferred (e.g. with a self-sample collection device with virus inactivation) | standard respiratory sample precautions | standard respiratory sample precautions | standard respiratory sample precautions | standard respiratory sample precautions | standard respiratory sample precautions | standard respiratory sample precautions | standard respiratory sample precautions | standard respiratory sample precautions | standard respiratory sample precautions | standard respiratory sample precautions
Quality control | internal control (for sample flow/migration); positive and negative controls sold separately; | internal control; lyophilized positive control and negative control provided in the kit; | internal control; pos/neg control | internal control; pos/neg control | internal control; pos/neg control | internal control | internal control; pos/neg control | internal control; pos/neg control | internal control; pos/neg control | internal control | internal control; pos/neg control | internal control; pos/neg controls separate
Remote connectivity capacity | not required for reader independent tests; if device-based: remote export of data possible | test is compatible with readers and other data capture devices; internal memory to store results even if power cut and with the ability to report to country health information management systems using on onboard unique identifier or other personal data protection safetguard linking the test to the user. | no | n/a | optional | Smartphone compatible | n/a | n/a | n/a | n/a | n/a | n/a
Need for additional equipment | handheld or on desktop (<1kg); battery or solar power operated; >8hr rechargable battery life | no additional equipment required, but if required, it should have potential for digital connectivity through smartphone applications | no | Sofia / Sofia 2 instrument | LumiraDx Instrument (portable) | Smartphone | MD-Bio Heater (or thermal cycler) | CFX 96 PCR system | CFX 96 PCR system | Yes, fluorescece analyzer | Sofia / Sofia 2 instrument | Standard F Analyzer
Need for maintenance/spare parts | none, swap out or replace ancillary device when needed | none required as device-free | n/a | n/a | n/a | n/a | n/a | n/a | n/a | none required | n/a | n/a
Waste/disposal requirements | routine biohazard waste | not applicable | biohazard waste | biohazard waste | biohazard waste | General waste | biohazard waste | biohazard waste | biohazard waste | biohazard waste | biohazard waste | biohazard waste
Manufacturing | ISO 13485: 2016 compliant | WHO or stringent regulatory authority emergency use listing/authorization or WHO prequalification or other stringent regulatory authority approval | FDA EUA | FDA EUA | FDA EUA | FDA EUA | FDA EUA | FDA EUA | FDA EUA | CE | FDA EUA | CE
Accessibility | to maximize accessiblity, all product manufacturers should have the capability to rapidly scale-up production and offer the product at a cost that allows borad use, including in low- and middle-income countries. |  | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a
Note on K16: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Product info sheet: 91.72%
Note on L16: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Reference not available
Note on M16: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Reference not available Reply: Sensitivity: 85% [78.3-90.2] specificity: 98.9% [98.3-99.4]
Note on K17: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Product info sheet: 100%

## Indirect POCT
WHO TPP: Point of care test (POCT) for prior SARS-CoV-2 infection
Published: 28 Sept 2020
| Wolrd Health Organization v.1.0 (20 Sept 2020) |  | FDA EUA Selected Antibody POCT | [9 empty] | FIND Selected Antibody POCT
| Acceptable | Desirable | COVID-19 IgG/IgM Rapid Test | Wantai SARS-CoV-2 Ab Rapid Test | RightSign COVID-19 IgG/IgM Rapid Test | LYHER COVID-19 IgM/IgG Combo Test | healgen COVID-19 IgG/IgM Rapid Test | Orawell IgM/IgG Rapid Test | Rapid COVID-19 IgM/IgG Combo Test | COVID-19 IgM/IgG Antibody Detection Kit | SGTi-flex COVID-19 IgG | SARS-CoV-2 IgG/IgM Rapid Test | COVID-19 IgM/IgG Gold | AllTest COVID-19 IgG/IgM Rapid Test | COVID-19 IgM/IgG Rapid Test | Healgen COVID-19 IgM/IgG Rapid Test | COVID-19 IgM/IgG Rapid Test | Wondfo SARS-CoV-2 Ab Test | AMP Rapid SARS-CoV-2 IgM/IgG Test | SARS-CoV-2 IgM/IgG Rapid Test | NG IgG/IgM Rapid Test
Manufacturer |  |  | Assure Tech. (Hangzhou) Co. Ltd. | Beijing Wantai Biological Pharmacy Enterprise Co., Ltd. | Hangzhou Biotest Biotech Co., Ltd. | Hangzhou Laihe Biotech Co., Ltd | Verify Diagnostics | Jiangsu Well Biotech Co., Ltd. | Megna Health Inc. | Nirmidas Biotech | Sugentech, Inc. | Xiamen Biotime Biotechnology Co., Ltd. | Technogenetics srl | Hangzhou Alltest Biotech Co., Ltd. | VivaChek Biotech (Hangzhou) Co., Ltd. | Verify Diagnostics | BioMedomics, Inc. | Guangzhou Wondfo Biotech Co., Ltd., China | AMEDA Labordiagnostik GmbH | MEDsan GmbH | NG Biotech Laboratoires
Country |  |  | China | China | China | China | USA | China | USA | USA | Korea | China | Italy | China | China | USA | USA | China | Austria | Germany | France
URL |  |  | http://www.diareagent.com/contact.html | https://www.fda.gov/media/140030/download | https://www.fda.gov/media/138657/download | https://www.fda.gov/media/139409/download, http://www.lyher.com/default.aspx?pageid=14&pageType=detail&id=54 | https://verifydiagnostics.com/healgen-covid-19-igg-igm-rapid-test-receives-fda-eua-approval/ | https://www.fda.gov/media/142422/download | https://www.megnahealth.com/products/#rapid-antibody-test | https://www.nirmidas.com/nirmidas-covid19-igmigg-antibody-detection-kit-eua-authorized | https://sugentech.com/products/products-view.php?ct=7&target=32%27 | https://www.fda.gov/media/140443/download | https://www.technogenetics.it/it/prodotti/covid-19/rapid-test/rapid-test-sars-cov-2-igm-igg-gold | http://www.alltests.com.cn/EN/ProductMain.asp?cclassid=2&classid=51 | https://www.vivachek.com/en/prods/prod-rapidtest.html | https://verifydiagnostics.com/healgen-covid-19-igg-igm-rapid-test-receives-fda-eua-approval/ | https://www.biomedomics.com/products/infectious-disease/covid-19-rt/ | https://www.bilcare.com/SARS-CoV-2%20Antibody%20Test%20(Lateral%20Flow%20Method).pdf | https://www.amp-med.com/news-rapid-test-sars-cov-2 | https://www.medsan.eu/en/detection/ | https://ngbiotech.com/2020/07/13/ng-test-igg-igm-covid-19-obtains-trod-status/
Intended Use | easy to use test that can be used at the POC to detect prior SARS-CoV-2 infection |  | POC in patient care setting; Labs; not home use | professional use only; not home use | professional use only; not home use | professional use only; not home use | professional use only; not home use | professional use only; not home use | professional use only; not home use | professional use only; not home use | professional use only; not home use | professional use only; not home use |  | for professionals only; not home use | for professionals only; not home use | for professionals only; not home use | for professionals only; not home use | n/a | for professionals only; not home use | for professionals only; not home use | for professionals only; not home use
Target Population / patient | General population invovled in surveys; groups at high risk of exposure to SARS-CoV2 |  | general population | general population | general population | general population | general population | general population | general population | general population | general population | general population |  | general population | general population | general population | general population | general population | general population | general population | general population
Target use setting | outside labs; health facilities; community settings; by health care workers or lab techs | same as acceptable + can be performed by trained lay workers | trained lab personnel; not home use. | trained lab personnel; not home use. | trained lab personnel; not home use. | trained lab personnel; not home use. | trained lab personnel; not home use. | trained lab personnel; not home use. | trained lab personnel; not home use. | trained lab personnel; not home use. | trained lab personnel; not home use. | trained lab personnel; not home use. |  | trained lab personnel; not home use. | trained lab personnel; not home use. | trained lab personnel; not home use. | trained lab personnel; not home use. | n/a | trained lab personnel; not home use. | trained lab personnel; not home use. | trained lab personnel; not home use.
Target molecule | at least one isotype; other biomarkers specific to prior infection | n/a | NP, Spike | Spike | spike RBD domain | S1 spike domain | S1 spike domain | spike RBD domain | NP | S1 & RBD domains of Spike | NP & Spike RBD | spike |  | n/a | n/a | S1 spike domain | n/a | n/a | n/a | n/a | n/a
Analytical sensitivity / LOD | currently no standard/units to express LOD; min detectable concentration of analyte | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a |  | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a
Analytical specificity | detect only SARS-CoV-2 specific isotype or biomarker | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a |  | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a
Sensitivity (PPA) | >90% | >95% | 100% (88.7-100) | 100% (88.7-100) | 100% (88.7-100) | 100% (88.7-100) | 100% (88.7-100) | 100% (93.8-100) | 100% (88.7-100) | 96.6% (88.3-99.1) | 96.7% (83.3-99.4) | 100% (88.7-100) | 99.2% (95.4-100) | 98.1% (89.9-100) | 97.3% | 97% (82.8-100) | 97% (82.8-100) | 96.2% (87-99.5) | 95.7% (90.8-98.4) | 95.1% (83.5-99.4) | 96.9% (93.9-98.6)
Specificity (NPA) | >97% | >99% | 98.8% (93.3-99.8) | 98.8% (93.3-99.8) | 100% (95.4-100) | 98.8% (93.3-99.8) | 97.5% (91.3-99.3) | 94.8% (88.5-97.8) | 95% (87.8-98) | 97.9% (92.8-99.4) | 100% (95.4-100) | 96.2% (89.5-98.7) | 98.3% (96.1-99.5) | 100% (83.9-100) | 100% | 97.5% (91.3-99.7) | 97.1% (90.1-99.7) | 100% (83.9-100) | 97.4% (86.2-99.9) | 98% (89.4-100) | 100% (92.9-100)
Type of analysis | qualitative | n/a | qualitative; lateral flow | qualitative; lateral flow | qualitative; lateral flow | qualitative; lateral flow | qualitative; lateral flow | qualitative; lateral flow | qualitative; lateral flow | qualitative; lateral flow | qualitative; lateral flow | qualitative; lateral flow |  | qualitative; lateral flow | qualitative; lateral flow | qualitative; lateral flow | qualitative; lateral flow | qualitative; lateral flow | qualitative; lateral flow | qualitative; lateral flow | qualitative; lateral flow
Interpretation | visual; portable automated reader | visual - device free | visual | visual | visual | visual | visual | visual | visual | visual | visual | visual |  | visual | visual | visual | visual | visual | visual | visual | visual
Sample type | plasma; serum; | n/a | whole blood; serum/plasma; fingerstick blood | whole blood; serum/plasma; fingerstick blood | whole blood; serum/plasma | serum/plasma | whole blood; serum/plasma; not fingerstick blood | whole blood; serum/plasma; not fingerstick blood | serum/plasma | serum/plasma | venous blood, serum, plasma (not figerstick) | venous blood, serum, plasma (not figerstick) |  | whole blood (fingertip/venous), serum, plasma | whole blood (fingertip/venous), serum, plasma | whole blood; serum/plasma; not fingerstick blood | whole blood (capilary, venous), serum, plasma | whole blood/serum/plasma | whole blood/serum/plasma | whole blood (fingerstick), serum, plasma | whole blood (fingerstick), serum, plasma
Test kit format | kit contains material for sample collection | same plus controls | intenal control; pos/neg control not included | intenal control; pos/neg control not included | intenal control; pos/neg control not included | intenal control; pos/neg control not included | intenal control; pos/neg control not included | intenal control; pos/neg control not included | intenal control; pos/neg control not included | intenal control; pos/neg control not included | intenal control; pos/neg control not included | intenal control; pos/neg control not included |  | intenal control; pos/neg control not included | n/a | intenal control; pos/neg control not included | intenal control; pos/neg control not included | n/a | n/a | intenal control; pos/neg control not included | intenal control; pos/neg control not included
End use profile | Health care facility worker | Lay person / non-professional staff | trained lab personnel | trained lab personnel | trained lab personnel | trained lab personnel | trained lab personnel | trained lab personnel | trained lab personnel | trained lab personnel | trained lab personnel | trained lab personnel |  | trained lab personnel | trained lab personnel | trained lab personnel | trained lab personnel | trained lab personnel | trained lab personnel | trained lab personnel | trained lab personnel
Training needs | <0.5 days | <2hr (quick reference) | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a |  |  | n/a | n/a | n/a | n/a | n/a | n/a | n/a
Sample preparation steps | 1 | 0 | 0 | 0 | 0 | 1 (separate plasma/serum) | 0 | 0 | 1 (separate plasma/serum) | 1 (separate plasma/serum) | 0 | 0 |  | 0 | 0 | 0 | 0 | n/a | 0 | 0 | 0
Reagents reconsitituion | if very simple to do | no, ready to use | ready to use | ready to use | ready to use | ready to use | ready to use | ready to use | ready to use | ready to use | ready to use | ready to use |  | ready to use | ready to use | ready to use | ready to use | n/a | ready to use | ready to use | ready to use
Need to transfer a precise volume of sample | if autofill or volume marking | no or a few drops | fill line | 10ul | fill line (10ul) | 10ul | fill line (5ul) | fill line (10ul) | fill line (10ul) | fill line (10ul) | 10ul | one drop (10-15ul) |  | 10ul (fill line) | 10ul (fill line) | fill line (5ul) | 15ul blood; 10ul serum/plasma | n/a | 10ul whole blood, 5ul serum/plasma | 10ul (fill line) | n/a
Number of timed steps | <2 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 1 | 1 | 1 | 1 | n/a | n/a | 1 | n/a
Time to result | <40min | <20min | 15min | 15min | 10min | 10min | 10min | 10min | 10min | 10min | 10min | 20min |  | 10min | 15min | 10min | 10min | 15min | 10min | 10min | 15min
Sample stability pre-testing | 30min RT (10-35oC); 8hr (2-8oC) | 3hr RT (10-35oC); 24hr (2-8oC) | whole blood 3 days at 2-8oC; serum/plasma 7 days 2-8oC | serum/plasma 7 day 2-8oC or 3 wks -15oC; whole blood used immediately | serum/plasma 3 days at 2-8oC; -20oC for longer-term; whole blood 2 days at 2-8oC; do not freeze blood | 2-8oC for 2 days; <-20oC long-term | serum/plasma 3 days at 2-8oC; -20oC for 1mo; whole blood 2 days at 2-8oC; do not freeze blood | serum/plasma 2-8oC for 5 days; -70oC for longterm; | serum/plasma <3 days at 2-8oC; -20oC for <3mo; -70oC longterm; | 4-8oC <3 days; -20oC <1 mo; | 2-8oC <5 days; -70oC longterm; | serum/plasma <14 days at 2-8oC; -20oC for <3mo; whole blood 2-8oC <3 days; |  | serum/plasma 7 days at 2-8oC; -20oC for long-term; whole blood 2 days at 2-8oC; do not freeze blood | n/a | serum/plasma 3 days at 2-8oC; -20oC for 1mo; whole blood 2 days at 2-8oC; do not freeze blood | anticoagulated blood 2-8oC for 5 days; serum/plasma 208oC 5 days or -20oC longterm; | n/a | n/a | serum/plasma 3 days at 2-8oC; -20oC for longterm; whole blood 2 days at 2-8oC; do not freeze blood | n/a
Result validity stability | fixed reading time | stored images or 6 weeks | 30min | 20min | 20min | 15min | 15min | 15min | 20min | 15min | 30min | 30min |  | 20min | 20min | 15min | 15min | n/a | 15min | 15min | n/a
Invalid rate | <5% invalid results with correct use | <1% invalid results with correct use | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a |  | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a
Operating conditions | 15-35oC; 20-80% RH; altitude 1500m | 10-40oC; 25-90% RH; altitude 3000m | room temperature | room temperature | room temperature | room temperature | room temperature | room temperature | room temperature | room temperature | room temperature | room temperature |  | room temperature | room temperature | room temperature | room temperature | room temperature | room temperature | room temperature | room temperature
Test kit stability and storage | 12mo, 2-30oC, 70% RH; | 18-24mo, 2-40oC, 90% RH | 2-30oC; | 2-30oC for 9mo | 2-30oC; | 2-30oC, dark | 2-30oC; | 2-8oC <9mo; | 4-30oC; | 4-25oC | 2-8oC | 5mo at RT; 9mo at 2-8oC |  | 2-30oC | 2-30oC | 2-30oC; | 2-30oC; | 2-30oC; | 2-30oC; | 2-30oC; | 2-30oC;
Stability of kit once opened | 1hr | 4hr | 1hr | 0.5hr | 1hr | 0.5hr | 1hr | 2-8oC <7 days | use immediately | 15min | 30 days at 2-8oC | use immediately |  | n/a | n/a | 1hr | 1hr | n/a | n/a | n/a | n/a
Specimen capacity & throughput | >5/hr per operator | >10/hr per operator | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a |  | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a
Safety precautions | standard blood collection safety precautions | n/a | standard | standard | standard | standard | standard | standard | standard | standard | standard | standard |  | standard | standard | standard | standard | n/a | standard | standard | n/a
Quality control | internal control (sample flow/migration); positive / negative controls sold separately | internal control; lyophilized positive/negative conrols provided | internal control; positive, negative controls optional | internal control; positive, negative controls optional | internal control; positive, negative controls optional | internal control; positive, negative controls optional | internal control; positive, negative controls optional | internal control; positive, negative controls optional | internal control; positive, negative controls optional | internal control; positive, negative controls optional | internal control; positive, negative controls optional | internal control; positive, negative controls optional |  | n/a | n/a | internal control; positive, negative controls optional | internal control; positive, negative controls optional | n/a | internal control; positive, negative controls optional | internal control; positive, negative controls optional | n/a
Remote connectivity capacity | device-based: remote export of data possible; no device: export through separate 3rd party reader | test is compatible with readers and other data capture devices | no | no | no | no | no | no | no | no | no | no |  | no | no | no | no | n/a | n/a | n/a | n/a
Need for additional equipment | handheld: <1kg; battery or solar power; >6hr rechargable battery life | no additional equipment | no | no | no | no | no | no | no | no | no | no |  | no | no | no | no | n/a | n/a | n/a | n/a
Need for maintenance/spare parts | none, swap out or replace device | none required as device free | no | no | no | no | no | no | no | no | no | no |  | no | no | no | no | n/a | n/a | n/a | n/a
Waste/disposal requirements | standard biohazardous waste disposal | smal environmental footprint; recyclable or compostable plastic | biohazsard waste | biohazsard waste | biohazsard waste | biohazsard waste | biohazsard waste | biohazsard waste | biohazsard waste | biohazsard waste | biohazsard waste | biohazsard waste |  | biohazsard waste | biohazsard waste | biohazsard waste | biohazsard waste | n/a | biohazsard waste | biohazsard waste | n/a
Accessibility | to maximize accessibility, all product manufacturers should have the capability to rapidly scale-up production and offer the product at a cost that allows broad use, including in low- and middle-income countries. |  | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a |  | n/a | n/a | n/a | n/a | n/a | n/a | n/a | n/a
Note on F15: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: For IgG/IgM combo test, combined sensitivity and specificity are reported