Supplementary Table S1: Detailed Site Verification Summary: September 2017-November 2018

Modules verified at each of thirteen sites are shown, with error codes highlighted.  The initial verification panels had a high number of volume errors, which led to operating procedure changes. A score of ≤4/6 was unacceptable and the module required retesting. Replaced modules were also tested before patient specimen testing commenced.

| Site/Panel | Module | Expected HIV VL | Difference (log cp/mL) | Score | Outcome |
| --- | --- | --- | --- | --- | --- |
| Site 1 | A1 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | 0.11 |  |  |
|  |  | High | 0.29 |  |  |
|  | A2 | Negative | Error2097 | 5/6 | Acceptable |
|  |  | Low | -0.10 |  |  |
|  |  | High | 0.08 |  |  |
|  | A3 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | -0.20 |  |  |
|  |  | High | 0.19 |  |  |
|  | A4 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | -0.20 |  |  |
|  |  | High | 0.31 |  |  |
| Site 2 | A1 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | -0.11 |  |  |
|  |  | High | 0.09 |  |  |
|  | A2 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | -0.08 |  |  |
|  |  | High | -0.04 |  |  |
|  | A3 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | 0.05 |  |  |
|  |  | High | 0.06 |  |  |
|  | A4 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | -0.20 |  |  |
|  |  | High | 0.91 |  |  |
| Module Replacement | A3 | Negative | 0.00 | 5/6 | Acceptable |
|  |  | Low | -0.05 |  |  |
|  |  | High | Invalid |  |  |
| Site 3 V17V11WA1 | A1 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | -0.07 |  |  |
|  |  | High | 0.27 |  |  |
|  | A2 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | 0.12 |  |  |
|  |  | High | -0.03 |  |  |
|  | A3 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | -0.21 |  |  |
|  |  | High | -0.01 |  |  |
|  | A4 | Negative | Error2097 | 5/6 | Acceptable |
|  |  | Low | -0.27 |  |  |
|  |  | High | 0.02 |  |  |
| Module Replacement | B3 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | -0.10 |  |  |
|  |  | High | -0.26 |  |  |
| Site 4 | A1 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | 0.33 |  |  |
|  |  | High | 0.25 |  |  |
|  | A2 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | -0.27 |  |  |
|  |  | High | 0.05 |  |  |
|  | A3 | Negative | 0.00 | 5/6 | Acceptable |
|  |  | Low | 0.15 |  |  |
|  |  | High | Error2005 |  |  |
|  | A4 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | -0.16 |  |  |
|  |  | High | 0.14 |  |  |
| Site 5 | A1 | Negative | 0.00 | 5/6 | Acceptable |
|  |  | Low | Error5017 |  |  |
|  |  | High | 0.19 |  |  |
|  | A2 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | 0.01 |  |  |
|  |  | High | -0.11 |  |  |
|  | A3 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | -0.22 |  |  |
|  |  | High | -0.01 |  |  |
|  | A4 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | -0.15 |  |  |
|  |  | High | 0.27 |  |  |
| Site 6 | A1 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | 0.01 |  |  |
|  |  | High | 0.20 |  |  |
|  | A2 | Negative | Error2096 | 5/6 | Acceptable |
|  |  | Low | -0.34 |  |  |
|  |  | High | 0.07 |  |  |
|  | A3 | Negative | Error2096 | 5/6 | Acceptable |
|  |  | Low | 0.05 |  |  |
|  |  | High | -0.13 |  |  |
|  | A4 | Negative | Error2096 | 5/6 | Acceptable |
|  |  | Low | -0.13 |  |  |
|  |  | High | 0.14 |  |  |
| Site 7 | A1 | Negative | 0.00 | 5/6 | Acceptable |
|  |  | Low | Error2005 |  |  |
|  |  | High | 0.04 |  |  |
|  | A2 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | -0.03 |  |  |
|  |  | High | -0.01 |  |  |
|  | A3 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | -0.03 |  |  |
|  |  | High | -0.08 |  |  |
|  | A4 | Negative | 0.00 | 5/6 | Acceptable |
|  |  | Low | Error5007 |  |  |
|  |  | High | 0.25 |  |  |
| Module Replacement | A4 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | -0.15 |  |  |
|  |  | High | -0.19 |  |  |
| Site 8 | A1 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | -0.20 |  |  |
|  |  | High | 0.01 |  |  |
|  | A2 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | 0.06 |  |  |
|  |  | High | 0.04 |  |  |
|  | A3 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | 0.07 |  |  |
|  |  | High | 0.16 |  |  |
|  | A4 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | -0.07 |  |  |
|  |  | High | 0.17 |  |  |
| Site 9 | A1 | Negative | Error2097 | 5/6 | Acceptable |
|  |  | Low | -0.25 |  |  |
|  |  | High | 0.41 |  |  |
|  | A2 | Negative | Error5017 | 5/6 | Acceptable |
|  |  | Low | -0.13 |  |  |
|  |  | High | 0.05 |  |  |
|  | A3 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | -0.06 |  |  |
|  |  | High | 0.07 |  |  |
|  | A4 | Negative | Invalid (5016) | 5/6 | Acceptable |
|  |  | Low | -0.34 |  |  |
|  |  | High | 0.27 |  |  |
| Site 10 | A1 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | -0.16 |  |  |
|  |  | High | 0.12 |  |  |
|  | A2 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | -0.32 |  |  |
|  |  | High | -0.03 |  |  |
|  | A3 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | 0.03 |  |  |
|  |  | High | 0.00 |  |  |
|  | A4 | Negative | Error2096 | 4/6 | Unacceptable |
|  |  | Low | -0.1 |  |  |
|  |  | High | Error5017 |  |  |
| Verification Repeat | A4 | Negative | 0.0 | 6/6 | Pass |
|  |  | Low | 0.1 |  |  |
|  |  | High | 0.1 |  |  |
| Site 11 | A1 | Negative | Error2096 | 5/6 | Acceptable |
|  |  | Low | -0.06 |  |  |
|  |  | High | 0.25 |  |  |
|  | A2 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | 0.13 |  |  |
|  |  | High | 0.03 |  |  |
|  | A3 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | 0.05 |  |  |
|  |  | High | -0.02 |  |  |
|  | A4 | Negative | Error2096 | 5/6 | Acceptable |
|  |  | Low | 0.17 |  |  |
|  |  | High | 0.20 |  |  |
| Site 12 | A1 | Negative | 0.00 | 5/6 | Acceptable |
|  |  | Low | Invalid (5016) |  |  |
|  |  | High | 0.27 |  |  |
|  | A2 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | -0.36 |  |  |
|  |  | High | 0.06 |  |  |
|  | A3 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | -0.14 |  |  |
|  |  | High | -0.20 |  |  |
|  | A4 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | -0.25 |  |  |
|  |  | High | 0.10 |  |  |
| Module Replacement | A4 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | -0.19 |  |  |
|  |  | High | -0.30 |  |  |
| Site 13 | A1 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | -0.34 |  |  |
|  |  | High | 0.16 |  |  |
|  | A2 | Negative | Error2097 | 5/6 | Acceptable |
|  |  | Low | -0.33 |  |  |
|  |  | High | 0.05 |  |  |
|  | A3 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | 0.08 |  |  |
|  |  | High | -0.02 |  |  |
|  | A4 | Negative | Error2096 | 5/6 | Acceptable |
|  |  | Low | 0.00 |  |  |
|  |  | High | 0.22 |  |  |
| Module Replacement | A1 | Negative | 0.00 | 6/6 | Pass |
|  |  | Low | -0.06 |  |  |
|  |  | High | -0.23 |  |  |

Supplementary Table S2: Detailed Site EQA Summary: September 2017-November 2018

EQA results from instruments tested at each of thirteen sites are shown, with error codes highlighted.  A score of ≤7/8 was unacceptable and it was recommended that EQA was repeated after root cause analysis and corrective actions.

| Site | Score | Outcome | Comment | Log Difference (copies/ml) |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Neg | Log 3.0 | Log 3.7 | Log 4.7 |
| EQA Round 1 2018 |  |  |  |  |  |  |  |
| 1 | 8/8 | Pass |  | 0 | -0.30 | 0.00 | 0.05 |
| 2 | 8/8 | Pass |  | 0 | -0.03 | 0.12 | 0.10 |
| 3 | 8/8 | Pass |  | 0 | 0.14 | 0.15 | 0.01 |
| 4 | 8/8 | Pass |  | 0 | -0.04 | 0.04 | 0.16 |
| 5 | 8/8 | Pass |  | 0 | -0.01 | 0.06 | 0.02 |
| 6 | 7/8 | Acceptable | Load volume error | 0 | ERR-2097 | 0.07 | -0.03 |
| 7 | 8/8 | Pass |  | 0 | 0.05 | 0.00 | 0.10 |
| 8 | 8/8 | Pass |  | 0 | 0.07 | 0.00 | -0.02 |
| 9 | 8/8 | Pass |  | 0 | -0.04 | 0.18 | 0.14 |
| 10 | 8/8 | Pass |  | 0 | -0.02 | 0.06 | 0.03 |
| 11 | 8/8 | Pass |  | 0 | 0.08 | 0.17 | 0.12 |
| 12 | 8/8 | Pass |  | 0 | -0.17 | 0.01 | -0.11 |
| 13 | 8/8 | Pass |  | 0 | 0.06 | 0.05 | 0.11 |
| Mean |  |  |  | 0 | -0.08 | 0.04 | -0.05 |
| Range |  |  |  | 0, 0 | -0.36, 0.08 | -0.07, 0.16 | -0.21, 0.06 |
| Standard Deviation |  |  |  | 0 | 0.11 | 0.07 | 0.08 |
| 1 | 8/8 | Pass |  | 0 | -0.05 | -0.02 | -0.01 |
| EQA Round 2 2018 |  |  |  |  |  |  |  |
| 2 | 8/8 | Pass |  | 0 | 0.10 | 0.14 | 0.15 |
| 3 | 8/8 | Pass |  | 0 | -0.10 | -0.02 | 0.01 |
| 4 | 8/8 | Pass |  | 0 | -0.07 | 0.06 | 0.05 |
| 5 | 7/8 | Acceptable | Loading error | 0 | -0.10 | 0.02 | ERR-2096 |
| 6 | 8/8 | Pass |  | 0 | -0.07 | 0.02 | 0.04 |
| 7 | 8/8 | Pass |  | 0 | 0.16 | 0.04 | 0.01 |
| 8 | 8/8 | Pass |  | 0 | -0.19 | -0.06 | 0.02 |
| 9 | 8/8 | Pass |  | 0 | -0.26 | -0.01 | -0.03 |
| 10 | 8/8 | Pass |  | 0 | -0.14 | 0.06 | 0.08 |
| 11 | 8/8 | Pass |  | 0 | -0.13 | -0.07 | 0.04 |
| 12 | 8/8 | Pass |  | 0 | 0.06 | 0.07 | 0.10 |
| 13 | 8/8 | Pass |  | 0 | -0.12 | 0.06 | 0.00 |
| Mean |  |  |  | 0 | -0.07 | 0.03 | 0.03 |
| Range |  |  |  | 0, 0 | -0.26, 0.16 | -0.07, 0.17 | -0.03, 0.15 |
| Standard Deviation |  |  |  | 0 | 0.12 | 0.07 | 0.05 |