Tables

| Material | 3D-printing Technology | Dental Applications |
| --- | --- | --- |
| Polymer-based: - Castable Resins - Hard Polymer - Clear Hard Polymers - Resin Composite Tooth Shade - Resin Composite Gingiva Elastic Shade - Waxes - Polyethylene - Polylactic Acid - Polycarbonate - Polysulfide - Polycaprolactone - ABS - PEEK* - PEKK* | Vat-polymerization SLA Vat-polymerization DLP Vat-polymerization CLIP Vat-polymerization LCD Material Jetting Powder-based Fusion SLS Material Extrusion FDM | Casts, casted metal frameworks, pressed lithium disilicate wax restorations, surgical diagnosis, surgical guides, occlusal devices, deprogrammers, silicone indices, custom trays, interim restorations, denture teeth, mock-up restorations, denture bases, bone analogs, orthodontic aligners. |
| Metal-based - Co-Cr Alloys - Titanium - Gold | Powder-based Fusion SLS Powder-based Fusion SLM Powder-based Fusion EBM | Surgical guides, splinting frameworks for complete arch impression techniques, frameworks for removable partial dentures, frameworks for tooth- and implant-supported prostheses, crowns, dental implants, and maxillofacial prosthetic parts. |
| Ceramic-based - Zirconia* - Lithium Disilicate* - Hybrid Ceramics* | Vat-polymerization SLA Vat-polymerization DLP Material Jetting Material Extrusion FDM Powder-based Fusion SLS | Tooth-supported Restorations |

Table 1: Summary of the main applications and materials of 3D-printing in dentistry [1-10].

SLA= Stereolithography; DLP= Direct Light Processing; CLIP= Continuous Liquid Interface; LCD= Liquid Crystal Display; SLS= Selective Laser Sintering; FDM= Fused Deposition Molding; EBM= Electron Beam Melting; ABS= Poly(acrylonitrile/butadiene/styrene); PEEK= Polyether Ether Ketone; PEKK= Polyetherketoneketone; Co/Cr= Chrome Cobalt, *= Experimental Phase

| Trueness |  |  |  |
| --- | --- | --- | --- |
| Volumetric Changes | P-Value | Formlabs (Castable) | Shera (Provisional) |
| Veneers | 0.854 | 88 26 µm | 85 41 µm |
| Incisors Labial | 0.001 | -97 84 µm | -77 98 µm |
| Incisors Palatal |  | 22 83 µm | 64 91 µm |
| Molars | 0.002 | 53 19 µm | 77 42 µm |
| FPDs | 0.004 | 181 91 µm | 214 89 µm |
| Copings | p-Value | Formlabs (castable) | Shera (provisional) |
| Veneers | 0.909 | 18 7 µm | 17 2 µm |
| Incisors | 0.012 | 31 4 µm | 52 20 µm |
| Molars | 0.001 | 23 2 µm | 31 4 µm |
| FPDs # 14 | 0.001 | 47 9 µm | 52 6 µm |
| FPDs # 17 |  | 25 2 µm | 44 6 µm |

Table 2: Overall trueness values of the 3D-printed restorations.

| Precision |  |  |  |
| --- | --- | --- | --- |
| External dimensional changes | p-Value | Formlabs (LCL–UCL)* (castable) | Shera (LCL-UCL)* (provisional) |
| Veneers | 0.054 | -3- 179 µm | -26- 198 µm |
| Incisors labial | 0.892 | 59- 139 µm | 69- 160 µm |
| Incisors palatal |  | 59- 137 µm | 64- 149 µm |
| Molars | ≦ 0.001 | 2- 103 µm | -32- 188 µm |
| FPDs | 0.101 | 169 – 270 µm | 206 – 328 µm |
| Internal dimensional changes | p-Value | Formlabs (LCL–UCL)* (castable) | Shera (LCL-UCL)* (provisional) |
| Veneers | 0.002 | 0.5 – 36 µm | 7 – 28 µm |
| Incisors | ≦ 0.001 | 2 – 4 µm | 1 – 11 µm |
| Molars | 0.305 | 1 -5 µm | 2- 8 µm |
| FDPs # 14 | 0.012 | 6 – 12 µm | 3 – 13 µm |
| FDPs # 17 |  | 1 – 6 µm | 4 – 14 µm |

* LCL= Lower control limit; UCL: Upper control limit.

Table 3. Overall precision values of the 3D-printed restorations.